## laser_info
Laser Mfg | Laser Model | Wavelengths | Laser Type | Eyewear Lens Compatible | Ease In | Rec1 | Rec2 | Rec3 | VLT | Optical Density | Ease In Ranges | Website
Biolase | Waterlase iPlus | 2780nm | Erbium | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Biolase | Waterlase Express | 2780nm | Erbium | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
DeKa | SmartPerio | 1064nm | Nd:Yag | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Fotona | LightWalker AT | 1064nm / 2940nm | Nd:Yag / Erbium | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Fotona | LightWalker AT S | 1064nm / 2940nm | Nd:Yag / Erbium | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Fotona | Lightwalker ST-E | 2940nm | Nd:Yag | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Fotona | SkyPulse Endo | 2940nm | Nd:Yag | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Fotona | SkyPulse Versa | 2940nm | Nd:Yag | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Fotona | NightLase | 2940nm | Nd:Yag | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Millennium | PerioLase MVP-7 | 1064nm | Nd:YAG | Pi23 |  | 701.Pi23 | 757.Pi23 | 626.Pi23.Pi17 | 70% | 870-950nm (OD 4), 950-1000nm (OD 5), 1000-1550nm (OD 7), 1550-2750nm (OD 4), 2800-11000nm (OD 6)
Denmat | NVPro3 | 808nm (OD 4) | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
AMD | Picasso+ | 800-820nm OD 5 | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
AMD | Picasso+ Lite | 800-820nm OD 5 | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
AMD | Clarion | 800-820nm OD 5 | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
Brassler | BLU Dental Microlaser | 810nm | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
CAO | Precise LTM | 810nm | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
Clincians Choice | Clincians Choice | 808nm | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
Denmat | Sol | 810nm | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
THOR | 810nm Infra Red Single Laser Dental Probe | 810nm OD 3 | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
THOR | 810nm 1W laser cluster (5 x 200mW) | 810nm OD 4 | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
THOR | 810nm 200mW single laser probe | 810nm OD 3 | Diode Laser | Pi1 |  | 701.Pi1 | 757.Pi1 | 626.Pi23.Pi17 | 70% | 800-830nm (OD 5), 9000-11000nm (OD 7)
Zolar | Photon dental diode laser | 810nm OD 5+/980nm OD 5+ | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Biolase | Epic X | 940nm +/- 10nm | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm(OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Biolase | Epic Pro | 940nm +/- 20nm, 980nm +/- 20nm | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Biolase | Epic Hygiene | 980nm +/- 10nm | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
King Medical | Beamer STL 980 | 980nm | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Sirona Dentsply | Siro Blue | 655-665nm OD1, 960-985nm OD5 | Blue Light Laser, Diode | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Sirona Dentsply | Siro Blue Plus | 440-450nm OD5, 655-665nm OD1, 960-985nm OD5 | Blue Light Laser, Diode | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Ultradent | Gemini Laser | 810nm / 980nm | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Zolar | Photon Plus Dental Diode Laser | 810nm OD 5+/980nm OD 5+ | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Ivoclar Vivadent | Odyssey Diode | 810nm +/- 20nm | Diode Laser | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)
Convergent | Solea Laser | 9,300nm | Co2 | Pi19 |  | 701.Pi19 | 757.Pi19 | 626.Pi14 | 40% | 520-532nm (OD 0.49), 9000-11000nm (OD 7)
DeKa | Smart US-20D CO2 Laser | 10,600nm | Co2 | Pi19 |  | 701.Pi19 | 757.Pi19 | 626.Pi14 | 40% | 520-532nm (OD 0.49), 9000-11000nm (OD 7)
DeKa | SmartXide Ultraspeed2 CO2 Laser | 10,600nm | Co2 | Pi19 |  | 701.Pi19 | 757.Pi19 | 626.Pi14 | 40% | 520-532nm (OD 0.49), 9000-11000nm (OD 7)
GPT Dental | DENTA CO2 | 10600nm | Co2 | Pi19 |  | 701.Pi19 | 757.Pi19 | 626.Pi14 | 40% | 520-532nm (OD 0.49), 9000-11000nm (OD 7)
Lightscalpel | LS-1005 | 10600nm | Co2 | Pi19 |  | 701.Pi19 | 757.Pi19 | 626.Pi14 | 40% | 520-532nm (OD 0.49), 9000-11000nm (OD 7)
Lightscalpel | LS-2010 | 10600nm | Co2 | Pi19 |  | 701.Pi19 | 757.Pi19 | 626.Pi14 | 40% | 520-532nm (OD 0.49), 9000-11000nm (OD 7)
Zap Lasers | Softlase | 630-670nm (OD 0.7), 800-980nm (OD 3.6) | Dental | Pi17 |  | 701.Pi17 | 757.Pi17 | 626.Pi23.Pi17 | 34% | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7)

## Loupe_types
Mfg | Mod | Size | Insert Part Number
Orascoptic | 01 Frame | 50-19 (on the inside of the nose bridge) | 01S
Orascoptic | 01 Frame | 53-19 (on the inside of the nose bridge) | 01M
Orascoptic | 01 Frame | 55-19 (on the inside of the nose bridge) | 01M
Orascoptic | 01 Frame | 58-19 (on the inside of the nose bridge) | 01L
Orascoptic | 02 Frame | 49-19 (on the inside of the nose bridge) | 02S
Orascoptic | 02 Frame | 51-19 (on the inside of the nose bridge) | 02S
Orascoptic | 02 Frame | 53-19 (on the inside of the nose bridge) | 02L
Orascoptic | 02 Frame | 54-19 (on the inside of the nose bridge) | 02L
Orascoptic | Victory | 53-19 Victory Small (located on the nose bridge) | VM
Orascoptic | Victory | 55-19 Victory Medium (located on the nose bridge) | VM
Orascoptic | Victory | 57-19 Victory Large (located on the nose bridge) | VL
Orascoptic | Rudy Project | Rudy/Rydon Standard lens | IVR
Orascoptic | Rudy Project | Rudy/Rydon Deep B lens | IVL
Orascoptic | Rudy Project | Rave | IVR
Orascoptic | Revolution | Revolution | IVR
Orascoptic | XV1 | XV1 | IVL
Orascoptic | Legend | Small 52-36 | LM
Orascoptic | Legend | Medium 55-38 | LM
Orascoptic | Legend | Large 58-40 | LM
Orascoptic | Legend | 52-44 | IVL
Orascoptic | Tempo | Small | IVR
Orascoptic | Tempo | Large | Primo
Orascoptic | Dragonfly | One size | IVL
Surgitel |  | 46-18 | 02XS
Surgitel |  | 50-18 | 02S
Surgitel |  | 50-20 | 02M
Surgitel | Aero | 50-20 | IVL
Surgitel |  | 50-24 | 52
Surgitel | Aero | 53-19 | VM
Surgitel | Aero | 52-20 | 01L
Surgitel |  | 53-21 | 01S
Surgitel | Oakley Flak Jacket | One size | IVL
Surgitel | Ergo | S/48-17 | LM
Surgitel | Ergo | L/ 55-19 | SR
Surgitel | Sundog | One size | SR
Surgitel | Contour flex | One size | IVL
Zeiss | Standard | 50-18 | VS
Zeiss | Standard | 54-18 | 01L
Zeiss | Standard | 53-20 | 01S
Zeiss | Standard | 56-18 | VL
Q-Optics | Eclipse | 131 (Petite) | IVR
Q-Optics | Eclipse | 139 (Medium) | IVL
Q-Optics | Eclipse | 150 (Large) | IVL
Q-Optics | 877 | 54-19 | LM
Q-Optics | 877 | 56-19 | LM
Q-Optics | 879 | 51-19 | LM
Q-Optics | 879 | 53-19 | LM
Q-Optics | Prism | One size | IVL
Perioptix | Advista | One size | IVR
Perioptix | Rimz | One size | IVL
Perioptix | Hoagie | One size | IVL
Sheer Vision | Under Armour Enforcer | One size | IVL
Sheer Vision | Under Armour Igniter | One size | IVL
Sheer Vision | Titanium Medium | One size | 01S
Sheer Vision | Zone XL | One size | IVL
Sheer Vision | Surge | One size | IVR
Sheer Vision | Performance Oval Medium | One size | 02S
Sheer Vision | Del Rey | One size | IVR
Ultra Light Optics | Flat | One size | 02L
Ultra Light Optics | Sport/TC | One size | IVL
Designs for Vision | Yeoman | 54-22 (inside frame, upper left) | 52
Designs for Vision | Yeoman | 52-24 (inside frame, upper left) | 52
Designs for Vision | Yeoman | 52-22 (inside frame, upper left) | 52
Designs for Vision | Yeoman | 50-24 (inside frame, upper left) | 52
Designs for Vision | Yeoman | 50-22 (inside frame, upper left) | 52
Designs for Vision | Yeoman | 48-24 (inside frame, upper left) | 50
Designs for Vision | Yeoman | 48-22 (inside frame, upper left) | 50
Designs for Vision | Yeoman | 48-20 (inside frame, upper left) | 50
Designs for Vision | Yeoman | 46-24 (inside frame, upper left) | 50
Designs for Vision | Yeoman | 46-22 (inside frame, upper left) | 48
Designs for Vision | Yeoman | 46-20 (inside frame, upper left) | 48
Designs for Vision | Yeoman | 44-22 (inside frame, upper left) | 48
Designs for Vision | Yeoman | 42-22 (inside frame, upper left) | 48
Designs for Vision | Yeoman | 55-17 (inside frame, upper left) | 01S
Designs for Vision | Yeoman | 56-18 (inside frame, upper left) | VL
Designs for Vision | Steam | 52-16 | IVL
Designs for Vision | 077 48-18 (Oval) | One size | 02XS
Designs for Vision | 047 51-18 (Round) | One size | 48
Designs for Vision | 062 54-19 | One size | 01M
Designs for Vision | 109 52-17 | One size | VM
Designs for Vision | 052 57-17 (Aviator) | One size | 52
Designs for Vision | Steam | One size | IVL
Designs for Vision | Icon | One size | IVL
Designs for Vision | Nike Retro | One size | IVR
Designs for Vision | Nike Run | One size | SR
Designs for Vision | DVI Sport Small | One size | IVR
Designs for Vision | DVI Sport Large | One size | IVL
Designs for Vision | Nike Skylon Ace | One size | SR
Designs for Vision | Skyview | One size | IVR
Lumadent | Argon | One size | IVL
Lumadent | Argon 2 | Traditional | IVL
Lumadent | ClariTi 2 | Traditional | IVL
Lumadent | InfiniTi 2 | Traditional | IVL
Lumadent | Stellar | Traditional | IVL
Lumadent | Argon 2 | Ergo | IVL.R
Lumadent | ClariTi 2 | Ergo | IVL.R
Lumadent | InfiniTi 2 | Ergo | IVL.R
Lumadent | Stellar | Ergo | IVL.R
Lumadent | Standard | One size | Primo
Univet | Ash | Small | IVR
Univet | Ash | Large | LM
Univet | Techne | One size | IVL
Univet | Model 546 | One size | IVL
Univet | Iris | One size | IVR
Heine | S-Frame | Sport frame/wraparound | IVL
Andau | Bolle | One size | IVR
Andau | ProGear | One size | IVL
Andau | MOS | One size | IVL.R
Andau | MOS-W | One size | IVL.R
Andau | Blues LG | One size | IVL.R
Andau | Blues MD | One size | IVL.R
Andau | Indie LG | One size | IVL.R
Andau | Indie MD | One size | IVL.R
Andau | Soul LG | One size | IVL.R
Andau | Soul MD | One size | IVL.R
Andau | Jazz | One size | IVR
Xenosys | IPAL-001-21-10 | One size | IVR.R
Xenosys | IPHX-001-21-10 | One size | IVL.R
Xenosys | IPAX-001-21-04 | One size | IVL.R
Orascoptic Current Models | Tempo | Ergo | Primo
Orascoptic Current Models | Tempo | Traditional | Primo
Orascoptic Current Models | Triumph | Ergo | IVL.R
Orascoptic Current Models | Triumph | Traditional | IVL.R
Orascoptic Current Models | Phantom | Traditional | IVL
Orascoptic Current Models | Phantom | Ergo | IVL.R
Orascoptic Current Models | Elevate | Traditional | IVL.R
Orascoptic Current Models | Elevate | Ergo | IVL.R
Orascoptic Current Models | Dragonfly | Traditional | IVL
Orascoptic Current Models | Dragonfly | Ergo | IVL.R

## Lens_details
Wavelength | Lens | Summary | OD | VLT | Shape | Material | Website | Image | Graph | Price(from)
| Gi2 | Fair functional VLT, High precision identification of red and blue objects, minimally fatiguing to wearer although relatively heavy | 574-595nm (OD 6), 596nm (OD 5.5), 736-755nm (OD 7) | 0.18 | flat | Glass | https://innovativeoptics.com/gi2-laser-glasses-frame/ | https://innovativeoptics.com/wp-content/uploads/2021/12/Gi2-circle-lenses.jpg | graphs/Gi2_plot.png | 395
| Pi3 | Significant orange to red light transmission, lightweight, relatively inexpensive | 190-532nm (OD 7), 850-900nm (OD 5), 900-1070nm (OD 6), 960-1064nm (OD 7), 9000-11000nm (OD 7) | 0.26 | curved | Polycarbonate | https://innovativeoptics.com/pi3-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/05/Pi3-laser-glasses-NEW.jpg | graphs/Pi3_plot.png | 195
| Gi2.Gi1 | Fair functional VLT, high precision identification of red and blue objects, minimally fatiguing to wearer although relatively heavy | 574-595nm (OD 6), 596nm (OD 5.5), 736-755nm (OD 7), 900-925nm (OD 4), 925-1000nm (OD 5), 1000-1025nm (OD 6), 1025-1400nm (OD 7), 1400-2400nm (OD 5), 2770-10600nm (OD 4), 10600nm (OD 5) | 0.15 | flat | Glass | https://innovativeoptics.com/contact/ | https://innovativeoptics.com/wp-content/uploads/2021/12/Gi2-circle-lenses.jpg | www/graphs/Gi2.Gi1_plot.png
| GiT1 | Fair VLT | 532nm (OD 7), 755nm (OD 8), 800-900nm (OD 5), 900-925nm (OD 4), 925-1000nm (OD 5), 1000-1025nm (OD 6), 1025-1400nm (OD 7), 1400-2400nm (OD 4), 2770-10600nm (OD 5) | 0.3 | flat | Glass | https://innovativeoptics.com/git1-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/12/GiT1-circle-lenses.jpg | graphs/GiT1_plot.png | 295
| GiT5 | Above average VLT | 650-655nm (OD3), 655-666nm (OD4), 666-675nm (OD5), 675-1025nm (OD6), 755nm (OD8), 1025-1400nm (OD7), 1400-2400nm (OD5), 2770-10,600nm (OD4), 10,600nm (OD5) | 0.6 | flat | Glass | https://innovativeoptics.com/git5-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/12/GiT5-circle-lenses.jpg | graphs/GiT5_plot.png | 395
| Gi1 | High functional VLT, high precision identification of red and blue objects, minimally fatiguing to wearer although relatively heavy | 900-925nm (OD 4), 925-1000nm (OD 5), 1000-1025nm (OD 6), 1025-1400nm (OD 7), 1400-2400nm (OD 5), 2770-10600nm (OD 4), 10600nm (OD 5) | 0.65 | curved | Glass | https://innovativeoptics.com/gi1-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/07/Gi1-laser-glasses.jpg | graphs/Gi1_plot.png | 295
| Pi4 | Fair transmission of green light, lightweight and relatively inexpensive | 190-450nm (OD 4), 650-665nm (OD 2), 730-1085nm (OD 5), 755nm (OD 7), 1064nm (OD 7), 9000-11000nm (OD 7) | 0.39 | curved | Polycarbonate | https://innovativeoptics.com/pi4-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/05/Pi4-laser-glasses-NEW.jpg | graphs/Pi4_plot.png | 189
| Pi14 | Low visible transmission of mostly green light, lightweight and intended for IPL only | 300-1400nm (OD 1), 9000-11000nm (OD 7) | 0.09 | curved | Polycarbonate | https://innovativeoptics.com/shade-3-ipl-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2019/11/8x2-IPL-Shade3.png | graphs/Pi14_plot.png | 59
| Pi11 | High VLT, lightweight and relatively inexpensive | 190-370 (OD 5), 9000-11000nm (OD 7) | 0.9 | curved | Polycarbonate | https://innovativeoptics.com/pi11-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/11/Pi10-laser-glasses.jpg | graphs/Pi11_plot.png | 59
| Pi23 | Moderate transmission of light across the visible spectrum, blocks near infrared, lightweight | 532nm (OD0.5) 1010-1090nm (OD7) 1064nm (OD10) 1650-2700nm (OD4) 2700-3000nm (OD5) 9000-11000nm (OD7) | 0.25 | curved | Polycarbonate | https://innovativeoptics.com/pi23-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2022/02/Pi19-laser-glasses.jpeg | graphs/Pi23_plot.png
| Pi10 | High VLT, lightweight and relatively inexpensive | 2780-3000nm (OD 4), 9000-11000nm (OD 7) | 0.9 | curved | Polycarbonate | https://innovativeoptics.com/pi10-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/11/Pi10-laser-glasses.jpg | graphs/Pi10_plot.png | 129
| Pi7 | Low functional VLT, lightweight and great for patients | 581-599nm (OD 6), 583-595nm (OD 7), 585-589 (OD 8), 640-670nm (OD 5), 650-665nm (OD 6), 9000-11000nm (OD 7) | 0.15 | curved | Polycarbonate | https://innovativeoptics.com/pi7-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/10/Pi7-laser-glasses.jpg | graphs/Pi7_plot.png | 265
| Pi22 | Low VLT, lightweight and great for clinical observers | 585-603nm (OD5), 587-600nm (OD6), 595nm (OD8), 900-1080nm (OD5), 970-1070nm (OD 6), 9000-11000nm (OD 7) | 0.09 | curved | Polycarbonate | https://innovativeoptics.com/pi22-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/04/PP26-laser-glasses.jpg | graphs/Pi22_plot.png
| Pi18 | Low VLT, lightweight and relatively inexpensive | 190-537nm (OD 8), 650-670nm (OD 2), 720-770nm (OD 7), 755nm (OD 8), 770-800nm (OD 6), 800-900nm (OD 5), 900-1077nm (OD 6), 9000-11000nm (OD 7) | 0.13 | curved | Polycarbonate | https://innovativeoptics.com/pi18-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/12/Pi18-laser-glasses.jpg | graphs/Pi18_plot.png | 195
| Pi16 | Fair VLT, lightweight and relatively inexpensive | 200-400nm (OD 4), 730-840nm (OD 5), 750-800nm (OD 7), 9000-11000nm (OD 7) | 0.32 | curved | Polycarbonate | https://innovativeoptics.com/pi16-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/04/PP26-laser-glasses.jpg | graphs/Pi16_plot.png | 195
| Pi8 | High VLT, lightweight and relatively inexpensive | 1010-1091nm (OD 5), 1020-1078nm (OD 6), 1030-1080nm (OD 7), 9000-11000nm (OD 7) | 0.5 | curved | Polycarbonate | https://innovativeoptics.com/pi8-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/10/Pi8-laser-glasses.jpg | graphs/Pi8_plot.png | 195
| Pi5 | Low VLT, lightweight and great for patients | 190-470nm (OD 5), 600-800nm (OD 1), 800-1750nm (OD 5), 830-1700nm (OD 6), 840-1600nm (OD 7), 870-1140nm (OD 8), 9000-11000nm (OD 7) | 0.09 | curved | Polycarbonate | https://innovativeoptics.com/pi5-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/04/Pi5-laser-glasses-2.jpg | graphs/Pi5_plot.png | 195
| GiT7 | Above average VLT | 430nm-500nm (OD 5), 500-515nm (OD 2), 515-532nm (OD 0.5), 655-1000nm (OD 5), 1000-1025nm (OD 6), 1025-1400nm (OD 7), 1400-2400nm (OD 5), 2770-10600nm (OD5) | 0.44 | flat | Glass | https://innovativeoptics.com/git7-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/12/GiT7-circle-lenses.jpg | graphs/GiT7_plot.png | 295
| Pi15 | Low VLT, lightweight and broadly protective | IPL Shade 5 300-1400nm (OD 1.5), 9000-11000nm (OD 7) | 0.07 | curved | Polycarbonate | https://innovativeoptics.com/shade-5-ipl-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2019/11/8x2-IPL-Shade5.png | graphs/Pi15_plot.png | 59
| Pi1 | High VLT, lightweight and relatively inexpensive | 800-830nm (OD 5), 9000-11000nm (OD 7) | 0.7 | curved | Polycarbonate | https://innovativeoptics.com/pi1-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/04/pi1.jpg | graphs/Pi1_plot.png | 175
| Pi19 | Grey, neutral density filters with above average VLT, made of lightweight materials to reduce fatigue | 520-532nm (OD 0.49), 9000-11000nm (OD 7) | 0.4 | curved | Polycarbonate | https://innovativeoptics.com/pi19-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2022/02/Pi19-laser-glasses.jpeg | graphs/Pi19_plot.png | 250
| Pi17 | Green lenses with high VLT, made with lightweight materials to reduce fatigue | 200-455nm (OD 5), 650-665nm (OD 1), 725-800nm (OD 5), 800-955nm (OD 5), 955-1080nm (OD 6), 9000-11000nm (OD 7) | 0.34 | curved | Polycarbonate | https://innovativeoptics.com/pi17-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2021/12/Pi17-laser-glasses.jpg | graphs/Pi17_plot.png | 209
| Pi30 | Violet-blue lens, made with lightweight materials to reduce fatigue | 610-630nm(OD 5), 630-660nm (OD6), 660-705nm (OD 8) | 0.19 | curved | Polycarbonate | https://innovativeoptics.com/pi30-laser-glasses-frames/ | https://innovativeoptics.com/wp-content/uploads/2018/10/Pi7-laser-glasses.jpg | graphs/Pi30_plot.png | 265
| PP19 | [7 empty] | graphs/PP19_plot.png
| Pi32 | Blue LIght (400-470nm) | 200-470nm (OD5), 630-685nm (OD3), 685-700nm (OD6), 700-730nm (OD 6), 730-1090nm (OD 7), 9000-11000nm (OD 7) |  |  | Polycarbonate |  |  |  | 189